## Slide 1: FLImBRUSH data processing

## Slide 2: Content
Definition of average lifetime
Determine FLImBRUSH acquisition window
  Simulation
  FLImBRUSH
Conditional value
Effect of sampling rates
FLImBRUSH system error/imperfections
How to determine alpha value for FLImBRUSH

## Slide 3: Average lifetime definition
Ideally, the integration is from 0 to infinity
Not possible for physical system
We have to chose a truncation length that minimize the error

## Slide 4: Simulation: single exponential decay
For single exp decay: the acquisition/truncation window should be at least 7 times the lifetime!
Q: True for multi-exponential decay?

## Slide 5: Simulation: double exponential decay
For multi-exponential data, can not use lifetime as a estimation of acquisition/truncation window!
Tau 1 = 3 ns, Tau 2 = 16 ns, a1 = 0.1, a2 = 0.9
Ratio = 103.9/15.73 =6.6
Tau 1 = 3 ns, Tau 2 = 16 ns, a1 = 0.9, a2 = 0.1
Ratio = 113.3/7.84 =14.45

## Slide 6: FLImBRUSH
acquisition length is determined based on:
  Data peak voltage: 1V
  Digitizer Effective Number Of Bit(ENOB) = 8 (ENOB is an estimation of digitizer circuit including noise and distortion, more info)
  Digitizer vertical range = 2 V.
  Signal decays to noise floor:
 | ENOB=7(manual) | ENOB=8 | From measurement data
Noise | 0.0156 | 0.0078 | 0.035

## Slide 7: truncation window for PpIX
Simulation setup:
16ns decay
Irf from FLImBRUSH
Assuming 20% before peak, 80% after peak

## Slide 8: FLImBRUSH acquisition window
Current FLImBRUSH acquisition:
Total: 500 points*0.4 ns/points = 200 ns.
Data after peak: 106 ns (17ns short for slowest iRF).
Solution: We can adjust the trigger delay by 50ns.

## Slide 9: Alpha vs Condition value
Sampling rate (GS/s) | Truncation length | Eps = 1.01 | Eps = 1.0001
20 | 3080 | 0.9837 | 0.9808
10 | 1540 | 0.9677 | 0.9619
5 | 770 | 0.9365 | 0.9252
2.5 | 385 | 0.877 | 0.856
We have been using 1.0001 for the past
Simulation:
Lifetime range: 0.3-20ns
Gaussian white noise: 55dB SNR (good approximation for FLImBRUSH)
Truncation window = 154ns (assuming PpIX)

## Slide 10: 5 GS/s

## Slide 11: 10 GS/s

## Slide 12: 20 GS/s

## Slide 13: Summary
The effect of condition value (1.01 vs 1.0001) is minimum. The biggest different is for lifetime > 16ns.
The conditional value just give us a more consistent/convenitent way to estimate alpha.
Under ideal case (0 dc offset, gaussian noise) the larger the alpha, the wider the range of lifetime we can fit.
We will use 1.01 for FLImBRUSH

## Slide 14: Effect of up sampling result from dye measurements
2 deconcolution setting:
  154ns window(PpIX) with its corresponding alpha
  154ns window(PpIX) with alpha corresponding to 54.4ns window (V4)
  54.4ns window(matching V4) with its corresponding alpha(see below)
Sampling rate (GS/s) | 154 ns window | 54.4 ns window
5 | 0.9252 | 0.8024
10 | 0.9619 | 0.8958
20 | 0.9808 | 0.9465
Conditional value is 1.0001 here.

## Slide 15: Simulation of lifetime range of V4: 12.5 GS/s, 54.4ns(680 points),alpha = 0.916, single exp decay
V4 is good for lifetime range of ~ 0.5 ns to 6 ns
Notes: Highlight consistency with FLImBRUSH

## Slide 16: FLImBRUSH: 5 GS/s 154 ns alpha = 0.9252

## Slide 17: FLImBRUSH: 5 GS/s 154 ns alpha = 0.8024

## Slide 18: FLImBRUSH: 5 GS/s 54.4 ns alpha = 0.8024

## Slide 19: 5 GS/s, 54.4 vs 154 ns, alpha = 0.8024
154ns lifetime – 54.4ns lifetime

## Slide 20: FLImBRUSH: 10 GS/s 154 ns alpha = 0.9619

## Slide 21: FLImBRUSH: 10 GS/s 154 ns alpha = 0.8958

## Slide 22: FLImBRUSH: 10 GS/s 54.4 ns alpha = 0.8958

## Slide 23: 10 GS/s, 54.4 vs 154 ns, alpha = 0.8958
154ns lifetime – 54.4ns lifetime

## Slide 24: FLImBRUSH: 20 GS/s 154 ns alpha = 0.9808

## Slide 25: FLImBRUSH: 20 GS/s 154 ns alpha = 0.9465

## Slide 26: FLImBRUSH: 20 GS/s 54.4 ns alpha = 0.9465

## Slide 27: 20 GS/s, 54.4 vs 154 ns, alpha = 0. 9465
154ns lifetime – 54.4ns lifetime

## Slide 28: 12.5GS/s C440

## Slide 29: iRF measurement
V4
DASPI for all channels
With fiber probe (~6m path)
540 nm band
Delay fibers plugged into Ch3 to get modal dispersion for each delay line
May not fully capture differences in modal dispersion due to launch conditions
FLImBRUSH
Laser reflection for Ch1 (without fiber probe)
DCS for Ch2 and DASPI for Ch3 (with fiber probe)
Modal dispersion in fiber probe not captured in Ch1 iRF (shorter iRF  over estimation of lifetime)

## Slide 30: Summary
Using maximum alpha (given by conditional value) would give us the largest range of lifetime value we can fit. However, it worsen the std of the data and caused a consistent over estimation of lifetime
No significant improvement of standard deviation using 20 GS/s compared to 10 GS/s
We will use 10 GS/s for FLImBRUSH
Two issues highlighted in red will be explained in the following slide

## Slide 31: Reconstructed Decays: 10 GS/s 154ns different alpha
The decay always deviate around 250 point(25ns) for all channels
The data have different length because the tail of the data sometime go to very small negative value less than ~-10e-7, which can not be displayed in log scale

## Slide 32
What if we truncate the decays to 250 points (25ns) and compute lifetime?
The lifetime is with in 0.1ns
So something in the tail cause the decay to drag long.

## Slide 33: Over estimation of lifetime caused by DC Simulation result
DC = 0.0003
DC = 0
0.0003 DC will cause a overestimation of lifetime by ~0.20 ns for lifetime 2-6 ns

## Slide 34: FLImBRUSH DC varies during scan
Subject031_20210804\2021_08_04_neuro031_04_S.mat
Scan duration: 1 min
Points color coded by time(blue is start of the scan, red is end of scan)
The DC value clearly changes as a function of gain and time (likely caused by digitizer temperature)

## Slide 35
FLImBRUSH DC estimation
Moving average: use DC from +- 500 waveform (~+-1 in time), and with in gain of +- 5
Moving average help with DC estimation.
Red is raw waveform, no averaging
Notes: Red is on raw waveforms, regenerate with 4 averaging, compare

## Slide 36
Red is averaged WF, 4 averaging
FLImBRUSH DC estimation
Moving average: use DC from +- 500 waveform (~+-1 in time), and with in gain of +- 5

## Slide 37
FLImBRUSH DC estimation with moving average

## Slide 38: Effect of moving average on lifetime
Moving average helps with std but not much

## Slide 39: Histogram of DC different for all waveform

## Slide 40: Estimate DC in front vs back
Front = 0.1337
Back = 0.1344

## Slide 41: Signal reflection in data
There is always a big of signal reflection in all channels, around 0.7% of the peak signal

## Slide 42: Simulate the reflection, no DC
The reflection of signal causes the overestimation.

## Slide 43: Data with reflection removed replaced value 305-350 with average of 325-350, likely lower

## Slide 44: Data with reflection removed
Lower lifetime in Ch1&3 likely due to the value used to replace reflection is a bit lower.
The higher std can be expected as shown in the following slides.

## Slide 45: Lifetime std vs alpha value simulation
Lifetime 3-20 ns
Truncation window: 154 ns
Signal SNR = 50 dB
Sampling rate: 10 GS/s
Alpha: 0.9619 vs 0.8958

## Slide 46: Lifetime std vs alpha value simulation
Using a larger alpha will increase the std of lifetime.

## Slide 47: Strategy for FLImBRUSH data processing
Step 1: Rough estimation of lifetime:
  154ns truncation window, alpha = 0.9619, with signal reflection, 0 dc offset
Lifetime will be over estimated by less than 0.1 ns.
Single exp

## Slide 48: Strategy for FLImBRUSH data processing
Step 2: Accurate estimation of lifetime based on Step 1 using different alpha values.
Lifetime 0.5-6 ns
This should cover most of the cases
Single exp

## Slide 49: Strategy for FLImBRUSH data processing
Step 2: Accurate estimation of lifetime based on Step 1 using different alpha values.
Lifetime 1-8 ns
At this point, it is just a trade off between lifetime range and std.
Single exp

## Slide 50: double exp simulation parameter
Tau1: 0.3-16 ns
Tau2: 16 ns(PpIX)
Photon ratio (short/long) = 9/1
No DC
No reflection
Laguerre order = 12
Solve linear system for given and

## Slide 51: Average lifetime vs tau 1
Average lifetime and tau1 are linearly related.

## Slide 52: Lifetime range
True lifetime didn’t tell much about tau 1

## Slide 53
Lifetime range vs tau1
The fitting is clearly bad short tau1, how about use lower alpha?

## Slide 54
Lifetime range with lower alpha
Use a lower alpha helps but
1. still not good enough if you got NADH
2. Consistent underestimation

## Slide 55
with even lower alpha

## Slide 56: What if we use more Laguerre functions
Use higher Laguerre order helps but still not good enough for tau1 of 0.5ns

## Slide 57: reduce alpha
Boom, that is much better

## Slide 58: reduce alpha
Boom, that is much better

## Slide 59: Photon ratio = 19, 5% long component
This is done with 0.9400, similar to 0.9475

## Slide 60: Photon ratio = 39, 2.5% long component

## Slide 61: Photon ratio = 49, 2% long component

## Slide 62: Photon ratio = 99, 1% long component

## Slide 63: Summary
Use small alpha and higher order to expand range

## Slide 64: Removing blip from data processing
What I did: Removed blip (540-700) point from contributing to deconvolution

## Slide 65
From slide 64, using a higher alpha:
A slight higher lifetime ~0.05 ns
Larger standard deviation.
reflection is contributing to residual.
reflection is ignored

## Slide 66: Simulation with data points 540-700 removed
This region has higher std because tail of decay falls in the 540-700 window and got ignored. See next slide.
We can potentially use a small ignore window to improve std (slide 66).

## Slide 67: Simulation data for lifetime 5-7 ns
The tail of data falls in the window that got ignored causing higher std.

## Slide 68: 540-640 ignore window

## Slide 69: DC front start vs DC from tail
So I am wrong early, DC from start and tail has no effect for Ch2&3, little effect for Ch1.

## Slide 70: DC front start vs DC from tail
So I am wrong early, DC from start and tail has no effect for Ch2&3, little effect for Ch1.

## Slide 71: Compared to V4
Alpha = 0.9616, Truncation = 154 ns
Alpha = 0.8958, Truncation = 154 ns

## Slide 72: Single exp simulation

## Slide 74
So the reflection in data has minimum effect on single exp simulation.

## Slide 75: Double exp simulation
So the reflection in data has minimum effect on double exp as well.